PRODUCT DATA SHEET

3D NEWTON/ 2013HMKZ0001

Guarantee

| PRODUCT NAME 3D NEWTON |  |
| --- | --- |
| Rating: 220-230V~, 50/60Hz, 400W Weight: 500kg |
| Serial Number |  | MGF. Date |  |
|  | Please read the user manual carefully before use |  |  |
| Manufacturer | HANMED Co., Ltd. #112, GBMC, 155-1, Nongso-Ri, Juchon-Myeon, Gimhae-Si, Gyeongsangnam-Do, Korea Tel: +82-55-331-0575, Fax: +82-55-331-0547 |

Date of sale :   19.4.2011           Sign :

PART 1

TECHNICAL PASSPORT

- SAFETY

1) Warning

Warning: Only a physical therapist, a spine doctor, or a licensed specialist can use the device. Otherwise, the device can be used only under the instructions of a licensed specialist.

1) Check thoroughly for any abnormality across the device.

2)When you find any abnormality in the device, stop the operation            immediately and ask the manufacturer to check after taking appropriate       measures.

3) When you find any abnormality in the device or unusual symptoms in the subject, take appropriate measures such as putting the subject on a flat place and stopping the device.

4)When a sudden blackout causes the device to stop during operation,         unlock the fastening device for pelvis fixing with more than one helper to extricate the subject from the machine safely.

5) Prohibit other people except the subject from going near the device.

6) Do not remodel the device arbitrarily.

7) When the device breaks down, contact an expert after marking                appropriately.

8) If the subject feels uncomfortable during operation, stop the operation and        have the subject see a doctor, and then try again.

9) During operation, do not touch or manipulate the power-supplying part or parts in operation.

10) Do not make the slope too steep at the beginning.

11) Do not apply external shock to the device during operation.

12) Install the main body of the device on a flat place.

13) Danger, warning - If any of the following problems arise, carefully select the most appropriate angle for the particular use.

① High blood pressure

② Eye problems such as glaucoma, diabetes-related diseases

③ Esophageal hiatal hernia

④ Weak bones (osteoporosis), any fracture that has just occurred and which has not been cured yet, acupuncture of bone marrow

⑤ Encephalosclerosis

⑥ Heart failure or cardiovascular function failure that has not been cured yet

⑦ Severely swollen joint

⑧ Pregnant women and elders

14) Prevention method

Warning : The machine needs to be examined regularly.

Cautious : Use extension cable and electricity protector for safety.

A) Do not expose the machine to rain or moisture

B) Do not disassemble or redesign the machine without authorized person of HANMED CO.,LTD

C) Do not connect 3D NEWTON to power before checking voltage

D) Do not use damaged wire

E) Do not put liquid or external substance to colsole

F) Do not remove the cover

G) Do not repair the machine without authorized person of HANMED CO.,LTD

- FUNCTION

This medical  instruments, a fixed brace, make your body reform straight.

It consists of the back of a chair, a acupressure, a massage roller, a pad, a lumbar rotary, a cylinder motor, a foot and leg fixing machine, a height regulator, and a safety device. Also there are a controller, a back support, a support board, and an arm fixing equipment that manage the structures mentioned above.

This product can practice the finger-pressure and massage the spinal muscles and pelvis. It can bend your lumbar and pelvis to the settled angle to help your pelvis exercise. It will keep your lumbar, thoracic vertebrae and cervical vertebral distances between the bones by using the angle device. That will be very useful to fix your body.

It have a position control system within also, and it can be operated by the electric motor system.

| 3D NEWTON |
| --- |
| Computer Control System |  | ①Contact the state of the switch, polarity, set the dial switch after checking the flow meter unit will verify that it works properly. ② Will determine whether the ground is perfectly clear ③ Access to all of the code will certainly be. 1) Check the power voltage.(This unit is dedicated to AC 220V.) 2) All of the access code is correct and complete the check. (※NOTICE : When unsafe or heolgeoul heat is the cause of the fire.) 3) Is connected to ground completely check. 4) And check the connection status of the switch. 5) Various control devices (See the screen) and it works Take a test drive by. (At this time there is a very loud noise, or at least to work should not be used, please contact the manufacturer representative.) |
| Program Operation System1 |  | Turn NEWTON, click the folder you are coming to initialize. After checking the contents of the screen in order to enter the patient New ①New Patients (Data) when entering Click the button and then enter the patient Add. ②Patients (Data) and then later to enter button Click Complete Revision ③If there are modifications to modify the patient (Data) to select and modify the input information and button click Delete ④Deleted when you delete the contents of the patient the patient (Data), select button click |
| Program Operation System1 |  | 1) As soon as it started, All of health conditions and medical treatment progress during remedy can be monitored at a look 2) It monitors supplement, complement, re-modification for medical treatment method during remedy in State Screen 3) For ideal medical treatment, it has complex multimedia functions, for example audio function plays even music in same time which is able to enhance comfort and muscle relaxation of patient |

3.SPECIFICATIONS

1) Rated voltage and Frequency : AC220V, 60Hz

2) Power Consumption : 400W

3) Muscles of the body against gravity, the device survives the exercise

4) Tilt angle feature

- Operating Range : Forward tilt range 0° - 60°

5) Body left and right rotating motor function

- right : 0~340° ±3°

- left : 0~340° ±3° Turn rate : 3~5rpm

6) Part is not a fixed feature

- Semi-automatic: Fixed.Speed ​​control when releasing the,

Fixation, suspension rates: 28 sec ±5 sec

- Fully automatic: Automatically turn off when turned off,

Release rate: 14 sec ±3sec

7) Pelvic fixture adjustment up and down

- Max : 0~300mm ± 5mm

8) Equipment weight : Body – 420kg / Control box – 80kg

4. COMPLETE SET

| Object | Manufacturer /Trademark | Type/Model | Technical Data | Standard | Mark(s) of Conformity |
| --- | --- | --- | --- | --- | --- |
| Power cord set | Hwajin KDK | KKP-4819R | 250V~,16A, H05VV-F, 3G0.75㎟ | IEC 60884-1 | VDE |
| Appliance inlet | Rong Feng | SS-120 | 250V~,10A | EN 60320-1, EN 60127-6 | VDE |
| Residual current operated circuit- breakers with integral over current protection | LS industrial system | 32KGRa | In:15A, Vn:220V~, Icn:15mA, 2protected poles | IEC 61009-1, IEC 61009-2-1 | KEMA |
| Circuit breaker for over current | LS industrial System | BS32a | In:15A, Vn:220V~, 2protected poles | IEC 60898 | KEMA |
| Noisefilter | Fine Suntronix | SN-M3H-CM | 250V~,3A | EN 133200 | TUV |
| Relay | Honeywell | SZR-LYseries | 240V, 24Vd.c. | EN 50081 | TUV |

| Power supply | SYNQOR INC | AQ0300MU | Input:100-240V~5A Output:24 Vd.c. | UL 60601-1 | UL |
| --- | --- | --- | --- | --- | --- |
| DC motor | SPG | S9D Series | 24V d.c., 120W | EN 60601-1 | CE |
| Actuator | HIWIN | LAS3 | 24V d.c. | EN 60601-1 | CE |
| Linear Actuator | LINAK | LA31 | 24V d.c, | EN 60601-1 | TUV |

5.TECHNICAL MAINTENANCE

- How to maintain

- Clean surface of table with non-toxic bacteria cleanser.

- Clean pelvic cushion using bacteria cleanser.

- Remove dust on the machine using vacuum once a week.

2) How to maintain the machine

- Check the following items once a week

- Inspect the machine by checking setup window

- Inspect to see if any wire is damaged

- Inspect connection of control BOX

- Inspect joint part

- Inspect pelvic fixing part

6. TROUBLESHOOTING LIST

- Inspect the above items when an error occurs

- Inspect the above items at least 1~2 times a month

- Inspect the above items by function when operating motor

- #5~8 Motor function is not operated

3D NEWTON Service Maintenance/ Cleaning Log

Maintain an accurate record of all service calls, visits and routine maintenance performed on machine. Present this log to hanmed service technician upon arrival; failure to maintain this log could void your warranty.

| Service Date | Technician Servicing | Service Order # | Technician Notes |
| --- | --- | --- | --- |

| Service Date | Technician Servicing | Service Order # | Technician Notes |
| --- | --- | --- | --- |

7.WARRANTY CONDITIONS

This product was created after various tests and experiments undertaken by our company technicians. Our company prides itself in complete product management and the meeting of high quality standards. However, in the case of any manufacturer's defect or normal malfunction found during the warranty period, please bring your product warranty to the place of purchase or to company headquarters for repair. ◈ Warranty ◈ ㆍFree - Free warranty service on the product for 18 months ㆍCharge - Failure caused by natural disaster - Failure caused by inappropriate use - Failure caused by inappropriate handling - After 18 months has passed ㆍWarranty - Repair, exchange, refund policies are based on official notification of Ministry of finance and economy Product Fixing equipment Model 3D NEWTON Date of purchase Warranty Period 1.5 year from date of purchase Name Phone Address Manufacture date Purchase store Location Refer to user's manual before applying for service ㆍService : Purchase store or (055)331-0575 (excluding Holidays) ㆍe-mail : kinetrac@naver.com (sign)